## Slide 1: MPPCs for nEXO
Fabrice Retière for EXO photo-detector group

## Slide 2: The EXO concept
Charge readout
  2D position
  Charge arrival time
  Energy
Light readout
  Timing resolution
  Event start time
    No need for good energy resolution
  Energy

## Slide 3: Light/charge fluctuations
Single site data with 228Th source
Critical to detect simultaneously
Ionization electrons
Scintillation photons at 175nm

## Slide 4: EXO-200 looking for 0nbb
Nature 510, 229–234 (12 June 2014)
Background reduction:
  Energy resolution
  Low radioactivity
  LXe self shielding an dMulti-site / single-site (does not affect light)

## Slide 5: EXO-200

## Slide 6: Light detection in EXO-200

## Slide 7: EXO-200 resolution limited by APD noise
7
Average APD
Noise at WIPP
Expected APD noise after electronics upgrade
Active area of nEXO going up by more than x10
Electronics noise is going to get much worse...

## Slide 8: EXO-200 to nEXO
nEXO: at the conceptual stage
EXO-200 “operating” detector
Two drift regions (central cathode)
Charge collection on anode wires
Light readout by ~500 APDs

## Slide 9: nEXO baseline configuration Up to 4 m2 of photo-detectors
Photo-detector
Field rings
Vessel
Drawing by T.Tolba (Bern)
~130cm diameter
~130cm long (single drift space)
Cathode
Anode

## Slide 10: Scintillation photon detection requirements for nEXO
Work led by G.Cao (IHEP)
ENCq= electronics noise on charge
l = correlated avalanche (AP+CT)
DN = dark noise rate
Neglect light electronics noise

## Slide 11: Side issues about E-Field Charge may land on MPPCs
Photo-detectors
Shaping ring
Cathode (bulge)

## Slide 12: Aside about index of refraction
nLXe ~ nSiO2 ~ 1.7
Re(nSi) ~ 0.7-0.8
  Surprisingly large uncertainty
Very significant reflections at Lxe/SiO2-Si
  Anti-reflective coating with MgF2?
Reflected photons have a fair chance of being detected later on
  Account for reflectivity in simulations
Need careful scaling from measurements in gas/vacuum
http://refractiveindex.info/
http://www.ioffe.ru/SVA/NSM/Semicond/Si/optic.html

## Slide 13: Photo-detector specifications for nEXO
Parameters | Value
Photo-detection efficiency at 175-178nm (without AR coating measured in gas/vacuum) | ≥15%
Radiopurity: contribution of photo-detectors to the overall background | <1%
Dark noise rate at -100°C | ≤50Hz/mm2
Average number of correlated avalanches per parent avalanche at -100°C | ≤0.2
Single photo-detector active area | ≥1cm2
Capacitance | <50pF/mm2

## Slide 14: Assessing nuisance parameters TRIUMF test setup
Control temperature down to -110C
Trigger on dark noise pulses with threshold ~0.5PE
  Measure cross-talk as
    P1 = 1-N1/Ntrig
    N1 = number of single PE
Measure all pulses following trigger
  Fit when pulses are close to each others
  Use scope time for late pules
  Built time difference distribution between trigger and next pulse

## Slide 15: Amplitude and time after an avalanche
KETEK
25ns recovery
900ns recovery
Cross-talk
Still a bit of hole diffusion?

## Slide 16: Timing analysis at -100°C
Add recovery plot
Improve this plot
Scope dead time
Dark noise edge
Hamamatsu
KETEK
Dark noise edge yields ~0.5Hz/mm2
But some not so clear issue with long
time constants
Short AP time constant ~2% probability
Poor fit.
Not exponential

## Slide 17: Efficiency measurement at Stanford
In vaccum
Light source: Xe gas scintillation
Liquid nitrogen cooler
Use reference Rxxx PMT for light yield calibration

## Slide 18: Efficiency measurement at Stanford

## Slide 19: Low radioactivity issue
Hoping to have the photo-detectors contribute less than 1% of total background
  238U and 232Th are <9mBq/kg (<0.73pg/g) and <43mBq/kg (<10pg/g)
  Impossible to measure by counting: 1kg of material yield 1 count in 30h.
Assaying either by neutron activation in a reactor or Inductively Coupled Plasma Mass spectrometry
  Measure content of stable isotopes and assume equilibrium
  Assaying must happen in parallel with device characterization
Low radioactivity packaging: looking at bounding Silicon chip on quartz plates

## Slide 20: Low radioactivity packaging
nEXO requirements are extremely low
  Much lower than dark matter experiments
Strategy:
  Use as little material as possible
    E.g. Mount MPPCs directly on Photo-detector Readout Board. I.e. Without carrier board. Make testing very cumbersome however
  Assay every single pieces of material
    EXO would like to assay raw materials (e.g wire bound, epoxy) before they get used in the final product
      Assay a sample of silicon before packaging
  Use Nuclear Activation for Enhanced sensitivity

## Slide 21: Low radioactivity material
Silicon is expected to be fine
  Though, may depend on details...
Wire bound. Little material. Probably ok
Epoxy. Little material. Probably ok
  Some Epoxies were found to be acceptable in EXO-200
    Electrically conductive epoxies have not be tested though
  Also possible issue with contamination of the liquid Xenon
Carrier board material. Significant amount of material (more than Silicon) so risky
  EXO currently investigating using quark or Saphire
Metal for electrical connection
  Some copper has very low radioactivity. Can be provided by nEXO
Solder for board to board connections
  Can be very radioactive and should be avoided as much as possible
Other possible materials: optical thin films, window in front of the MPPC will have to be assayed

## Slide 22: Large area readout issues
PD area~4m2
# channel~ 100-10,000
Channel area up to 20×20cm2
Common base option may work
  Single photon identification compromise with >10nF
Total power < 50W
MPPC
3x3 mm2
0.32 nF
MPPC
3x3 mm2 + 6.6 nF
~1.4×1.4cm2
MPPC
3x3 mm2 + 15 nF
~2×2cm2
G. Visser
G. Visser (Indiana U.)

## Slide 23: Possible readout configuration
Photo-detectors
  2×2cm2
Carrier card
  Holding photo-detector
    May not be needed
  2.2×2.2cm2
Support and connection board
  16 readout ASICs on the back side
  All connections on the back side
ASIC
Photo-detectors
Carrier cards
Support and
Connection
board

## Slide 24: Possible configuration 2
16 ASICs per photo-detector readout boards (PRB)
1 coincidence unit per PRB and possibly ADC and timing capabilities
Would require ~24 PRBs
  x2 segmentation in z
  x12 in azimuth
Total number of photo-detectors
~70cm
~35cm
PRB photo-detector side
Note that the numbers don’t fully add-up accounting
for nEXO’s current baseline geometry. But it does not
matter right now.

## Slide 25: MPPC configuration
Area requirement 4 to 5 cm2
  Based on 3x3 mm2 devices
    8x8 matrix = 2.4x2.4 = 5.76 cm2 (need ~10,000 in that case)
    7x7 matrix = 2.1x2.1 = 4.41 cm2
    6x6 matrix = 1.8x1.8 = 3.24 cm2
  Or possibly on different sizes
Desirable to connect all the MPPCs on the silicon to minimize connections
  The question is whether or not to have some MPPCs in series vs parallel
    One option: 4x4 matrix of 6x6mm2 (or 8x8 of 3x3mm2) making 4 groups of MPPCs connected in parallel (either 4 6x6mm2 or 16 3x3mm2) and then connect the 4 groups in series
    Or for better uniformity connect first in series 4 6x6 mm2 (or 16 3x3mm2) and then connect in parallel.

## Slide 26: Using MEG MPPC for development
Investigate series vs parallel solution using MEG MPPCs at TRIUMF
  PCB holding the MPPC must fit on the cold chuck
Analog board
Schematics: L. Fabris, ORNL
Layout: M. Constalbles, TRIUMF?
Cold board
1 MEG MPPC
Cold board
4 MEG MPPCs
Analog board
Schematics: L. Fabris, ORNL
Layout: M. Constalbles, TRIUMF?
Test performances with
1 6x6mm2
4 6x6mm2 in series
4 6x6mm2 in parallel
Test performances with
16 6x6mm2 in parallel
4 6x6mm2 in series then parallel
4 6x6mm2 in parallel then series

## Slide 27: The current landscape
Parameters | Spec. | FBK-2010 | HPK | KETEK | SensL
Over-voltage | N/A | 5V | 2.5V | 2.5V | 
PDE at 175nm (%) | >15% | 10S | 17S , 19M | 0 | 0
Dark noise rate at -100C (Hz/mm2) | <50 | 103 S | 0.5T | 4T | ?
Cross-talk probability |  | 0.06S | 0.34T | 0.1T | ?
Total after-pulsing rate at -100C |  | ? | 0.16T | 0.25T | ?
Xt+AP within 10 ms at -100C | <0.2 | ? | 0.38T | 0.15T | ?
Recovery time |  | ? | 25nsT | 900nsT | ?
Pulse rise time (Gaussian s) |  | ? | 1.6nsT | 0.5nsT | ?
Pulse fall time(s) (exp. constant) |  | ? | 30nsT | 6/300nsT | ?
Capacitance (pF/mm2) | <50 | 330 | 35 | 100 | 80 / 3.2
Thorium content (mBq/kg) | < 43 | <13 A | <42 103 | ? | <2 106 N
Uranium content (mBq/kg) | < 9 | <2.5 102 A | <100 103 | ? | <94 106 N
Potassium content (mBq/kg) | ? | <1.4 102 A | <43 103 | ? | <4.7 106 N
S Measured at Stanford, T Measured at TRIUMF, A measured at U.Alabama with neutron activation
M Measured by the MEG collaboration, N measured by the NEXT collaboration by Ge counting

## Slide 28: Summary
Hamamatsu MPPCs are very well suited for nEXO if 2 issues can be addressed
  Lower cross-talk with trenches. Easy?
  Find a way to achieve the required extremely low radioactivity... Hard?
Would be nice add-ons:
  Increase efficiency using anti-reflective coating. Possible?
  Lower capacitance per unit area. Possible?

## Slide 29: Time scale
Identify compelling photo-detector(s) by the end of 2016
Build a small scale (10x10x10cm3) demonstrator by early 2017
Down-select by DOE in 2016 or 2017
  Chose technology, i.e. Germanium or Xenon
Construction starting in 2017 or 2018

## Slide 30: Add-on

## Slide 31: Low capacitance SiPM concept
Motivations
  Lower electronics noise. Can live with lower gain
  Less gain: reduce correlated avalanche rate
    Maintain same peak current which to first order is defined by DV/Rquench. Peak current is critical for good timing
  Narrow pulse. Fall time ~ Rquench x Cpix
Idea
  “Point” contact diode
  Need electric field to drift carrier to point contact
  Impact ionization only occurs close to point contact
  Single carrier detection (e- prefered)
  Block carriers from substrate from reaching the drift field region

## Slide 32
SiO2
Quenching resistor
metal
p-
p++
substrate
p-
metal
n++
High field region

## Slide 33
p-
p++