May 27
Taking too long to read large hdf5 files.
This is probably due to the current method of searching for each waveform.
Learning hdf5 in order to solve this issue.
Processed data. Data doesn't look good. Baseline keeps shifting, causing the peak finder to fail. Also, dark noise is quite high even at -100C (radioactivity?).
Upgrading peak finder to calculate local baseline (one peak width) as it scans the histogram for peaks.
Paul will retake data.
Fabrice suggests that problem may be caused by one of the four MPPCs being bad. Will switch to a different one.
WF on scope looks much better. Taking full dataset.
Constructed box ot of alimunum to protect reference MPPC.
May 28
Data from yesterday still looks bad.

Created algorithm to count zeros using the minimum of the WF to see if the issue is with the data or analysis.
Point at 1.5 is strange. This may be a data issue. Above plot made with only 1000 wf. Doing full set now.
Setup github.
Data is greatly improved with more statistics. See May28*-100* for plots.
Biggest issue is now the processing time. Need to make hdf5 reader more efficient.
Need to prepare talk.

May 29
Worked on talk.
Processed data.

May 31
Processed data to be used in talk. Finished preparing talk.

June 1
Photo-detector meeting with talk.
Now need to address cross-talk.