Carl Rethmeier

nEXO photo-detector characterization log 2015
May 27
Taking too long to read large hdf5 files.
This is probably due to the current method of searching for each waveform.
Learning hdf5 in order to solve this issue.
Processed data. Data doesn't look good. Baseline keeps shifting, causing the peak finder to fail. Also, dark noise is quite high even at -100C (radioactivity?).
Upgrading peak finder to calculate local baseline (one peak width) as it scans the histogram for peaks.
Paul will retake data.
Fabrice suggests that problem may be caused by one of the four MPPCs being bad. Will switch to a different one.
WF on scope looks much better. Taking full dataset.
Constructed box ot of alimunum to protect reference MPPC.
May 28
Data from yesterday still looks bad.

Created algorithm to count zeros using the minimum of the WF to see if the issue is with the data or analysis.
Point at 1.5 is strange. This may be a data issue. Above plot made with only 1000 wf. Doing full set now.
Setup github.
Data is greatly improved with more statistics. See May28*-100* for plots.
Biggest issue is now the processing time. Need to make hdf5 reader more efficient.
Need to prepare talk.

May 29
Worked on talk.
Processed data.
May 31
Processed data to be used in talk. Finished preparing talk.

June 1
Photo-detector meeting with talk.
Now need to address cross-talk.
Met with Fabrice and got new objectives.
Now need to stagger hours to maximize data taking. Will work 16 hours/day in total.
12 SiPMs to be measured. Will measure <PE>, <DN>, and <avalanche>.

June 2
Showed Paul how code works.
Got cross-talk analysis working using ntuple that finds dark noise pulses that were triggered on in the data. This pulse is at the centre of the WF. Finding the histogram of minimum values works the best in this case, since the probability of having multiple pulses in this 200ns range is very low at -100C. The ntuple mst be analyzed manually, however.

More data is being taken by Paul.
Working on afterpulse. Will simply integrate WF after the dark noise pulse at the tigger point.
Processed Bottom_Top4 data. This is <PE> data for the MEG with proper OV.
Made crosstalk more automated, as the previous method of reading of the height spectrum was very tedious.
Initial problem with automation was too few bins.
Implemented afterpulse. Testing both afterpulse and crosstalk now.
June 3
Paul had air pump problems while taking data. He thought it was a leak. It is now working fine, but this prevented data from being taken.
My email to Paul:
“So now the noise is completely eliminated on both channels. It was a grounding problem for the shielding on the cables, as well as on the normalization holder. Regarding electrical noise, both channels look as clean as they would if the lamp was turned off.

I removed the 4th piece of paper. I think what you saw before was just a high level of crosstalk. I was getting way too little data even at high overvoltages with 4 pieces of paper. I checked with 3, and even at an OV of 4V, there is still the right amount of light.

Also, I think we might have a light leak. There seemed to be less dark noise registering on the normalization when I covered the box with the black fabric. So just to be sure, I think it's best if we cover the box AND turn the room lights out while taking data, provided no one else is in the room. If you like, you could quantify this tomorrow by taking a measurement with the normalization with the lights on, and another with the lights off. Because it does not have to be cooled, you could do this while the bottom detector is cooling, so it won't cost you any extra time.

Lastly, please make sure the waveforms have no vertical offset when taking data. This will make the processing much simpler.

Right now I'm taking data with that same detector. I'll see how much data I can take tonight. I might just stay longer since the noise level is so good right now.”

The data for coated SiPM2 was taken.
June 4
The noise returned when Paul installed the next detector. Need to find a reliable solution.
My email:
“At least we know what the problem is for the top detector. It's the grounding of the shield around the cable, as well as the grounding of the holder itself. If this grounding can be improved, I think we will eliminate the noise. My theory is that the noise simply appears the way it does because when something like the cable simply moves inside the box, it causes the connection between some points of the foil to break. If we can improve that connection, it may solve the problem. Maybe try using a larger price of foil around the wire.”

Paul took data for SiPM1

From Fabrice:
“The MPPC labeled 1-9 (or 10?) are old generation MPPCs coated with various wavelength shifters. Being old generation their intrinsic performance such as dark noise and correlated avalanche are not so interesting. There is no need to measure dark noise with good precision for those. Furthermore I would even argue that you don't need to run them very cold, just cold that dark noise is not an issue. So, consider running them only down to -20V. You need a measure of dark noise for correcting a possible contribution to the efficiency but this could be taken from the same data set.
The other devices on the other hand, VUV2, and VUV3 have to be studied in details at -100C. VUV3 is the most interesting most likely.”

My email after Paul left:
“As soon as I started the pump, the noise came back. I re-shielded the cable, and this seemed to work for a few minutes, but then it returned again. I still haven't managed to get rid of it. I think that cable needs to be replaced as soon as possible, as I don't think it's grounding properly. I talked to Daryl about it, and he said tomorrow we can get it fixed, so I think you should talk to him as soon as you can tomorrow.”

Managed to collect data for SiPM3 after much fighting with the setup. Used foil to shield the cable, and played with the grounding till it worked.
June 5
Paul has been unable to take data due to noise. Fabrice has seen it as well.

I managed to solve it by using the copper tape to completely enclose the top detector. Any signal would now have to penetrate the copper shield, which is grounded to the base. It is screwed down using the same screws that hold the SiPM holder. Resulting signal:

Took data for VUV3 up to 13.5 OV for both trigger on lamp and dark noise (2 data sets).
June 8
Processed data for VUV3. Automation is not reliable, so had to manually pull final results from minHist.

Worked on Paul's report for the rest of the day.
June 9
Filled LN2 dewers. Completed by 10am.
Worked on conference application.
Processed data for normalization detector. <PE> slope seems to be increasing, indicating it is being cooled:

Paul will characterize normalization detector tomorrow at different OV.

Finished MEG CT plot.

Paul is testing S10362-13-050C detector to ensure we are only dealing with UV.

MEG <PE> processed data is bad. I split both the light and dark region into chunks when I analyzed it. I did not do this for the VUV3. This doesn't make sense, as most of the light comes in the first few 100ns. Reprocessing the data without this region splitting. This was initially done to try to get better dark noise error bars.

Did <PE> for Coated1. This detector seems pretty poor. Still re-processing data for MEG.

MEG data re-processed. MEG curve is now much closer to VUV3.
June 10
Added <PE> data up to SiPM #6. The BV level on some of the data needs to be corrected. Exceptions are VUV3 and SiPM2 data, which look good.

Did <CT> for SiPM #1. BV level for MEG and Coated1 clearly needs correcting. Paul will find the correct BV.

Fixed code to allow it to properly process positive pulses.

Did <CT> up to Coated6. BV for all except Coated2 needs to be redone.
June 11
Added <PE> for Coated 7-9, and <CT> for Coated 7.
June 15
Did <PE> for third VUV3 run.
Processed normalization data.
Hard drive filled up. Made new partition and transferred data to it.
Got ready to take VUV3 dark noise data tomorrow.
June 16
Worked on normalization calibration.
June 17
Procedure for gain studies developed.
Did gain study on normalization detector.
June 18
Talked to Fabrice. “Grass” between peaks is due to afterpulsing.
Lack of correspondence of data runs must be addressed ASAP.
Showed Paul how to do analysis.
Did VUV3 gain study.
4th VUV3 run again shows lack of correspondence.
Problem found in my analysis. I was using a window size of 1.5us in the VUV3 analysis instead of 3us. This was only the case for the VUV3 analysis.
Need to do a systematic study of the setup to determine source of problem in terms of reproducibility.
June 19
Did gain study on coated SiPMs.
Consecutive data runs for VUV3 at -100C with room lights off are different. Nothing was changed in the setup – one run was done after the other. The problem is most likely due to changes in the light output from the lamp.
Will need to use normalization detector to apply correction.
Did characterization of norm. PE.
Got fillNtp working for pulse finding on entire waveform for normailization detector.
Started gain analysis of first VUV3 run.
June 22
Had photo-detector meeting. Need to be more prepared next time.
Performed gain study on VUV3 runs and normalization.
Wrote NormGain.cxx.
Applied correction from normailization detector.
June 23
Fabrice wants detailed talk for Thursday.
Worked on talk most of day.
Wrote project plan:

Project plan:
1. Make presentation for Fabrice
  A. Describe setup
    I. Box
    II. Lamp and trigger
    III. Beam splitter
    IV. Filtering
      a. Testing for UV
    V. Attenuation
    VI. N2
    VII. Cooling
    IIX. Detector holders
      a. Reference
    IX. Amplifiers
    X. Power supplies
    XI. Shielding
      a. Lamp section
        i. Radio waves
      b. Air gaps
      c. Conductive tape and paper on lid
      d. Bottom detector
      e. Top detector
      f. Cables
    XII. Grounding
      a. Detectors
      b. Shielding
  B. Describe data taking
    I. Lamp position and power
    II. Allignment
      a. Top detector
      b. Bottom detector
    III. Groinding and noise
    IV. Lid and blanket
    V. N2 and Cooling
    VI. Oscilloscope
    VII. Fetch
  C. Describe data and processing
    I. Waveforms
    II. HDF5
    III. efficiency.exe
      a. findMin
      b. Peak finder
      c. Smoothing
      d. minHist for <PE>
      e. ntp
        i. Ability to exclude pulses with width and charge using ntp
    IV. fillNtp.exe
      a. minHist for <CT> and gain
    V. NormGain.exe
      a. ntp for gain
  D. Describe analysis
    I. <PE>
      a. minHist integration
      b. Calculation of results
      c. Calculation of errors
    II. <CT>
      a. minHist integration
      b. Calculation of results
      c. Calculation of errors
    III. <DN>
      a. minHist integration
      b. Calculation of results
      c. Calculation of errors
    IV. Gain
      a. ntp fitting
        i. Auto minHist for normalization
        ii. describe reasons for ntp vs minHeight
      b. Fitting to get BV
  E. Describe results
    I. <PE> results to date
      a. Problem with resproducibility
      b. BV problem from non-constant PE separation
    II. <CT> results to date
      a. BV problem from non-constant PE separation
      b. Potential afterpulse issue at low OV
        i. Show "grass"
    III. <DN> results to date
      a. Time consuming and low priority
    IV. Gain results to date
      a. Show norm. det.
        i. Show table of formula for BV and slope problem
      b. Show other det.
  F. Summary
    I. Successes
      a. Noise ellimination
        i. Grounding
        ii. Shielding
      b. <PE>, <CT>, <DN>, and Gain almost ready for mass testing
    II. Ongoing work
      a. Continued improvements to code
      b. Study of setup to determine cause of changes between runs
    III. Unresolved issues
      a. <PE> fluctuates significantly between runs
        i. Studying setup to solve this
      b. Afterpulse
        i. Will incorporate fitting into code
      c. Non-constant PE separation
        i. Research needed to find cause
        ii. Currently using 1PE peak for BV measurements
    IV. Future work
      a. Solve issues just described
      b. Correct results
      c. Take data with Ketek and FBK detectors

2. Test reproducibility of setup
  A. Two consecutive runs with VUV3
    I. Cool to -100C
    II. Take 15 data points
      a. 0.5,1,1.5,2,3,4,5,6,...,12,13
    III. Repeat a second time
    IV. Measure <PE> and Gain for both norm. and VUV3
  B. If difference between runs, design next step based on observations
  C. If runs are the same, test for light leak
    I. If no light leak suspected, try bumping setup between runs
  D. If data is the same difference as before, characterize at same OV
    I. If this does not reveal anything, try less sensitive detector

3. Get correction working if necessary
  A. Program this while waiting for Paul
    I. Get Gain and <PE> trends for norm. det.
      a. Gain is automated
      b. Automate <PE> using similar method
    II. Apply correction

4. Test Coated detectors <PE> at same OV
  A. Find the OV of the Coateds using gain
  B. Take 2 runs at each point for statistics
  C. OV of 0.5-2.0

5. Compare VUV3 and MEG <PE>, <CT>, and <DN>
  A. 1 run each for <PE>
  B. Trigger on DN for <CT>
  C. Change temp for <DN> trend

6. Correct OV for old data using gain
  A. Can do this once reproducibility is tested
    I. May need to retake data

7. Develop code
  A. Make <PE> automated
  B. Automate <CT>

Collected all figures for presentation.
June 24
Worked more on presentation.
Interpreted data to try to fix reproducibility problem.
Runs separated by one day show change.
Runs taken concurrently show no change.
Runs taken consecutively show no change.
Blew dust off detector without removing it from holder and observed small increase in <PE>.
June 25
Improved alignment of detector, and data restored to peak values.
Alignment seems to be the problem.
Did presentation with Fabrice.
Retook data to confirm previous results.
Data matched previous run.
With Fabrice, removed beam splitter and lower collimator to deal with alignment.
VUV3 was very well aligned from previous run.
With Paul, marked shock using wire through collimator for future runs.
MEG detector next, but lots of dirt from ziploc.
June 26
Paul cleaned glass cover of MEG.
Multi-photon peak observed. Need to solve this problem, else a fourth filter may be necessary. The strange thing is that the VUV2 normalization does not have the same problem.
Added fourth filter. The problem came about because Paul removed the copper plate, which was partially blocking the light.
I am uneasy about the breakdown voltages.
July 1
Data taken for 4.5 hours with VUV3 on bottom and VUV2 on top. The lamp was turned on at the same time the first data point was taken. Everything was held constant during this run.
July 2
Results from previous day:

<PE> above; Orange = top; Blue = bottom

Gain above
July 7
Paul took data as on July 1 but without beam splitter:

This seems to indicate that the beam splitter is not the problem.
Jul 8
Sent July 7:

So for tomorrow, I would like to thoroughly understand what is going on inside the box. Before I left, I reinstalled the beam splitter, and left the lamp and N2 on.

-------------------------------------------------------------
I want to do 3 runs tomorrow.

Run 1: Take as much data as possible over the first 4 hours at work using both the top and bottom detectors. Leave the cooling pump off. This measurement should be done with both devices at room temperature.

Turn lamp off and have lunch for an hour.

Run 2: Same as run 1, except turn the lamp on at exactly the same time the first data is taken. Take data for 4 hours to see how the lamp changes.

Turn off lamp and remove beam splitter.
Wait 1 hour.

Run 3: This is the control. The procedure will be the same as was done today, except with more data points. Do this for 4 hours. This must be done in case something changes between today and tomorrow.
-------------------------------------------------------------

This is a 14 hour job at the very least. What do you think of splitting our time as before, with one of us arriving in the afternoon?

The data above was taken as planned.

Note: The lamp was turned off before RUN2 and RUN3 were started.
July 9
After analyzing RUN1-RUN3, I can conclude that the lamp is not the primary problem. There is still a problem of some kind of oscillation of the <PE> on the devices, but this is much less (factor of 7) than the change in <PE> level observed when the lower detector was cooled. The lack of cooling was the only difference between this data and the data taken on Canada Day.

Paul's data from July 7 is under the same conditions as RUN2, which doesn't make much sense.

The correction seems to be working, at least it did on yesterday's data:
The <PE> also changed linearly with temperature yesterday, while with the cooling, the curve was very non-linear. However, this is most likely due to the small scale of the change.

July 7 with correction:
The fact that the second and last point above have the same amplitude but such different <PE> indicates that the detector temperature is not the problem. However, it is conceivable that the amplifier may be part of the problem. The July 7 data seems to implicate the lamp, while yesterday's data points toward something else.
July 10.
The SiPM was left on overnight before RUN1 was done.

The beam splitter has been ruled out as the source of the problem. This leaves the SIPM, amplifiers, HV supply, amplifier power, pulse generator, lamp, and box geometry as the possible sources of the problem. The HV can be ruled out by analyzing the amplitude. The amplifier and amplifier power would not affect the efficiency of the detector, only the gain, ruling this out. The non-random nature of the trend found on July 7 rules out variations in box geometry, as the temperature was fairly constant.
This leaves the pulse gen., lamp, and SiPM. RUN2 and RUN3 seem to indicate the lamp is not the issue, as it was turned off before starting the runs.
The pulse gen. can be ruled out from the fact that the oscilloscope uses the same trigger. A fault in the cable to the lamp can be ruled out from the smooth shape of the <PE> curve in the July 1 and 7 data.

This indicates that the SiPM may be changing in efficiency over time. This seems unlikely, however.
To test this, we planned to do two tests:

Test1: Leave the entire apparatus except the N2 systems on for 2 hours. After the 2 hours, turn the top SiPM (VUV2) on and start collecting data for 2 hours.

Test2: Turn everything except the SiPM off and wait 2 hours. After 2 hours, turn everything back on and collect data for 2 hours.

When we attempted to start Test1, the flash lamp made inconsistent buzzing sounds, and the current fluctuated between the usual 0.09A and 0.06A. After switching between internal and external trigger, the lamp seemed to stabilize. However, while taking the data, I noticed it fluctuated. Analysis of this data resulted in:

When I went to start Test2, the lamp would not turn on with the external trigger. I switched it to internal trigger as before. The current fluctuated between 0.09A and 0.11A (0.11A is expected). After a few minutes it decreased to 0.07A. Switching between external and internal trigger had no effect. The lamp needs to be replaced.

This is not surprising based on the datasheet from Hamamatsu. The lamp is guaranteed to output 1e9 flashes. We trigger 1 flash/ms. The lamp therefore has a lifetime of 1,000,000s = 16,667min = 278hr = 35 eight hour days.

http://www.hamamatsu.com/resources/pdf/etd/5W_Xe-F_TLSZ1006E.pdf

The lamp does not explain the lack of reprodicible re
The only alternative explanation for the opposing <PE> trends in the July 1 data is that some amount of frost may have built up on the lower detector and obscured some of the light.
July 21
Spent most of the day correcting the BV. Smoothing at 30ns radius seems to work the best. Done all the coated devices. (3.5hr)
Also talked to Fabrice briefly.
Added variable smoothing to NormGain. Troubleshooted it for a while. The problem was that the smoothing took place at different times relative to the baseline calculation in Normgain and fillNtp.
Also worked out dark noise calculation method.
Added triggerTime to findCleanPulse.

Still need to do BV for MEG and VUV3. MEG may take a long time, as I suspect it has high density data.
July 22
To do:
Finish BV for VUV3 and MEG - done
Fix CT error bars - done
Do trigger time dark noise macro
Merge bitBucket code
Add CT for Coated 3, 8, and 9
Get CT for VUV2
Compare the dark noise rate to the afterpulse rate.

Updated NormGain and helped Paul with analysis.
Discussed peak finder with lloyd.
July 23
Paul will install new lamp and carry out tests 1 and 2.
Helped Paul more with analysis.
Talked to Lloyd about his code.

To do:
Do trigger time dark noise macro - done
Merge bitBucket code
Add CT for Coated 3, 8, and 9
Get CT for VUV2
Compare the dark noise rate to the afterpulse rate.
July 24
To do:
Merge bitBucket code - done
Add CT for Coated 3, 8, and 9
Get CT for VUV2
Plan presentation
Make code for Igor plot right.

Talked to Fabrice.
Need to make plots as on page 10 of Igor's publication.
Need to test lamp.
Fabrice suggested that the lamp arc may be changing. I think this would explain all the inconsistencies we were seeing before. The solution would be to make a collimator for the lamp.
Will need to test current lamp to see if the results from Canada Day can be reproduced.
July 27
To do:
Add CT for Coated 3, 8, and 9
Get CT for VUV2
Plan presentation
Make code for Igor plot right. - done
Make code for Igor plot left - done
July 28
Flights – emailed Fabrice
Fix legends on other plots
Improve next pulse amp plot – in progress
Need to be able to manipulate plot as ntuple
Add CT for Coated 3, 8, and 9
Get CT for VUV2