## Slide 1: Qualitative
Xenon flash lamp
Bandpass
filter wheel
Light tight enclosure
Vacuum enclosure
Use as Faraday cage
as well
Cold plate
SiPM and PCB
LN2
Regulator
PID controller
Temperature
sensor
Amplifier and bias
Sipm
signal
viewport
Rail
Vacuum
line
Rail
Minimal system
Allow to investigate the response of SiPM to 175nm light qualitatively
Reproducibility will be a serious issue (no independent light monitor)
In air, the flash lamp has to be very close to viewport. EMI concern.
Could tune UV light output by moving flash lamp but very sensitive to water content
Cold temperature operation should be ok
Obsolete

## Slide 2: Critical qualitative cross-check
Rail
Light tight enclosure
Vacuum enclosure
Use as Faraday cage
as well
Cold plate
SiPM and PCB
LN2
Regulator
PID controller
Temperature
sensor
Amplifier and bias
Sipm
signal
2 viewports
Teflon tube
Alpha source
Xenon gas
Gas enclosure
Critical cross-check
Verify sensitivity to Xenon scintillation light (alpha source in xenon gas)
Source light yield should be stable (provided Xenon density does not change)
No trigger and no control over light yield (only qualitative)
Vacuum
line
Obsolete

## Slide 3: Quantitative
Xenon flash lamp
Neutral density
filter wheel
Bandpass
filter wheel
Light tight enclosure
Vacuum enclosure
& Faraday cage
LN2
Regulator
PID controller
Amplifier and bias
viewport
Rail
Vacuum line
TEFLON based
Integrating
sphere
Calibrated PMT
High speed oscilloscope
Nitrogen purge
System accommodates
Xenon flash lamp
Visible light from optical fibers
Xenon scintillation chamber
Complete system
Use calibrated PMT to simultaneously measure light yield
Use integrating sphere to uniformize the light hitting the photo-detectors
Accommodate alpha source and laser light sources for cross-checks (375, 405nm)
Allow to purge the light tight enclosure with Nitrogen or Argon (low VUV absorption)
Obsolete

## Slide 4: Final
Xenon flash lamp
TEFLON based
Integrating
sphere
Calibrated PMT
Iris
Teflon tube
Alpha source
375, 405, 467, 637 nm
Laser (direct or through fiber)
Source 1: Xenon scintillation
Source 3: UV/visible fast light sources
Source 2: Xenon flash lamp
Xenon gas in/out
SHV
Optical
fiber
SHV
BNC
Viewport
Filter
slider
PPD
Nitrogen/O2
gas in/out
Faraday
cage wall
Manual
stage
BNC
Filter
slider
Cold
N2 gas
Cold plate
Iris
Iris
Filter slider
BNC
MPPC
BNC

## Slide 5: Flash lamp dimension (metric)

## Slide 6: PMT (R9875P) with base (E10679) dimension